ELECTRONIC SUPPLEMENTARY MATERIAL

Miniaturized paper-based analytical device for the portable analysis of phyto-cannabinoids in plant and oral fluids

Dymphy Houtzager, Sergio Armenta, José Manuel Herrero-Martínez, Héctor Martínez-Pérez-Cejuela*

Department of Analytical Chemistry, University of Valencia, Dr. Moliner 50, 46100, Burjassot, Valencia, Spain

*Corresponding authors:	Héctor Martínez-Pérez-Cejuela

Tel.: +34 9635 44062

e-mail: Hector.Martinez-Perez@uv.es

Table of contents

Page S1. GC-MS settings (Table S1)

Page S2. Qualitative assessment (Table S2)

Page S3. Interference studies for qualification (Table S3)

Page S4. Characteristics of used papers (Table S4)

Page S5. Interference studies for quantification (Table S5)

Page S6. Table comparison (Table S6)

Page S7. PAD features and advantages (Table S7)

Page S8. Pearson plot of phyto-cannabinoid quantification (Figure S1)

Table S1. GC-MS settings reference method.

| Parameter | Value |
| --- | --- |
| Carrier gas | Helium |
| Oven temperature (start) | 150 °C min-1 |
| First ramp | 10 °C min-1 to 250°C |
| Oven temperature (end) | 250 °C per 10 min |
| Flow | 1 mL min-1 |
| Inlet temperature | 250 °C |
| Inlet mode | Splitless |
| Solvent delay | 5 min |
| Ion source temperature | 276 °C |
| Quadrupole temperature | 150 °C |

Table S2. Cannabis samples with the most abundant cannabinoid present and corresponding color observed with the 4-AP color reagent.

| Cannabis sample | Most abundant cannabinoid | Observed colour1 |
| --- | --- | --- |
| 0001 | CBD | Pink |
| 0002 | CBD | Pink |
| 0006 | Mix CBD/Δ9-THC/CBN | Purple |
| 0012 | CBN | Blue |
| 0014 | CBN | Blue |
| 0017 | CBN | Blue |
| 0018 | Mix Δ8-THC/Δ9-THC/CBN | Blue |
| 0020 | CBN | Blue |
| 0023 | CBN | Blue |
| 0024 | Mix Δ8-THC/Δ9-THC/CBN | Blue |
| 0025 | CBN | Blue |
| 0026 | Mix Δ8-THC/Δ9-THC/CBN | Blue |
| 0029 | CBN/Δ9-THC | Blue |

1Blue color corresponds to THC/CBN rich, the pink color corresponds to CBD rich, any other color is inconclusive.

Table S3. Interferences from relevant compounds of cannabis with the 4-AP reagent using the developed PAD method. Reagent A & B volume 1 µL; sample volume 0.5 µL; reading time 5 min.

| Herb | Result1 |
| --- | --- |
| Salvia lavandulifolia | no color |
| Boldo (peumus boldus) | Blue color |
| English lavender (Lavandula angustifolia) | no color |
| Thymus hyemalis | other color |
| Sen hojas (cassia angusfolia) | no color |
| Illicium verum | other color |
| Papaver rhoeas | no color |
| Citrus aurantium | other color |
| Equisetum arvense | no color |
| Matricaria chamomilla | other color |
| Taraxacum officinale | other color |
| Red tea | other color |
| Green tea | other color |
| Chai tea | other color |
| Coffee | other color |

1Other color: Reaction produced a color different than pink and blue.

Table S4. Different properties of the Whatman® filter papers used during this work.

| Whatman number | Paper type | Pore size (µm) | Thickness (mm) |
| --- | --- | --- | --- |
| W1 | Qualitative | 11 | 0.18 |
| W3 | Qualitative | 6 | 0.39 |
| W4 | Qualitative | 20-25 | 0.21 |
| W5 | Qualitative | 2.5 | 0.20 |
| W41 | Ashless quantitative | 20-25 | 0.22 |
| W42 | Ashless quantitative | 2.5 | 0.20 |
| W50 | Hardened Low Ash | 2.7 | 0.12 |
| W54 | Hardened Low Ash | 20-25 | 0.19 |
| W113 | General Purpose and Wet-Strengthened | 30 | 0.42 |
| W114 | General Purpose and Wet-Strengthened | 25 | 0.19 |
| W541 | Hardened Ashless | 20-25 | 0.16 |
| W542 | Hardened Ashless | 2.7 | 0.15 |

Table S5. Interferences from relevant compounds of cannabis with the Fast Corinth V reagent using the developed PAD method. Drugs concentration 0.05 mg mL-1; reagent volume 1 µL; sample volume 0.5 µL; reading time 15 min.

| Component | Pseudo-absorbance | Concentration THC (mg mL-1) |
| --- | --- | --- |
| Drugs of interest |  |  |
| Nicotine | 0.004 | <LOD |
| Caffeine | 0.003 | <LOD |
| Codeine | 0.002 | <LOD |
| Diazepam | -0.003 | <LOD |
| Cocaine | 0.000 | <LOD |
| Solid samples of interest |  |  |
| Synthetic cannabinoid 1 | 0.006 | <LOD |
| Synthetic cannabinoid 2 | -0.008 | <LOD |
| Chai tea | 0.002 | <LOD |
| Coffee | -0.008 | <LOD |
| Tobacco | -0.005 | <LOD |
| Red tea | -0.002 | <LOD |
| Green tea | -0.003 | <LOD |

Table S6. Comparison between the developed PAD procedure and similar methods reported in the literature.

| Analytes | Detection | Support material | Sample matrix | Detection time (min)a | RSD (%) | LODs (mg L-1) | Ref. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Δ9-THC, CBD and CBN | Colorimetric | Polydimethylsiloxane | Cannabis plant | 1 | ≤9 | 3.7 | [9] |
| Δ9-THC and CBD | Electrochemical | Paper (wax printing) | Cannabis oil | 15 | <5 | 3 | [14] |
| Δ9-THC, CBD and CBN | Colorimetric | Aluminium thin layer – Silica | Cannabis plant | - | - | Not determined | [15] |
| Δ9-THC, CBD and CBN | Colorimetric | Paper (stationery material) | Oral fluids and cannabis plant | 15 | <15 | 3 | This work |

Table S7. PAD features and advantages from the use of the present method.

| PAD features1 |  |
| --- | --- |
| Low cost | < 0.1 Euro cents |
| Lightweight | < 5 g |
| Required materials | Laminating pouch, paper filter, laminator |
| Detection system | Scan/smartphone |
| Throughput | 75 samples per hour |
| Low reagent/sample consumption Qualification | 7.2 µg 4-AP 40.7 µL EtOH 0.12 µL HCl 2 M 0.72 mg NaOH 7.2 µL H2O 24 µL sample |
| Low reagent/sample consumption Quantification | 240 µg Fast Corinth V 24 µL H2O 12-24 µL sample |
| Type of samples | Liquids (direct) Solid (after extraction) |

1For a one PAD card (24 analysis positions)

Figure S1. Pearson plot of phyto-cannabinoid quantification using the developed PAD method and the GC-MS reference method; the concentration points are the average from 3 different replicates.